## Slide 1: OMA-CD-DWAPI-2015-0049 PCH State Diagram and Error Handling
Submitted To: CD (DWAPI)
Date: 29 November 2015
Availability: Public OMA Confidential
Contact: Max Hata, masato.hata.uf@s1.nttdocomo.com
Takafumi Yamazoe, yamazoet@nttdocomo.com
Source: NTT DOCOMO
X
USE OF THIS DOCUMENT BY NON-OMA MEMBERS IS SUBJECT TO ALL OF THE TERMS AND CONDITIONS OF THE USE AGREEMENT (located at http://www.openmobilealliance.org/UseAgreement.html) AND IF YOU HAVE NOT AGREED TO THE TERMS OF THE USE AGREEMENT, YOU DO NOT HAVE THE RIGHT TO USE, COPY OR DISTRIBUTE THIS DOCUMENT.
THIS DOCUMENT IS PROVIDED ON AN "AS IS" "AS AVAILABLE" AND "WITH ALL FAULTS" BASIS.
Intellectual Property Rights
Members and their Affiliates (collectively, "Members") agree to use their reasonable endeavours to inform timely the Open Mobile Alliance of Essential IPR as they become aware that the Essential IPR is related to the prepared or published Specification. This obligation does not imply an obligation on Members to conduct IPR searches. This duty is contained in the Open Mobile Alliance application form to which each Member's attention is drawn. Members shall submit to the General Manager of Operations of OMA the IPR Statement and the IPR Licensing Declaration. These forms are available from OMA or online at the OMA website at www.openmobilealliance.org.

## Slide 2: Outline
The current PCH specs don't specify the way to report measurement state to apps
When apps send an one-shot request, if the Plug-In don't have any measurement values, the Plug-In is supposed to return an error to the application
But the app is not able to know the reasons why they received errors
This INPUT shows the measurement state transition and propose a framework for the error handling for apps
Explains the background for :
OMA-CD-DWAPI-2015-0048-CR_7.3_Behaviours_of_Plug_Ins_for_reporting_measurements_to_applications

## Slide 3: Measurement states
All types of PCH devices are always in one of the states as follows:
1) DISCONNECTED
  Plug-In is not connect to any devices
2) READY
  Plug-In is connected to the targeted device, and measurement is ready to start.
3) MEASURING
  Measurement is in progress, but no measurement result has been returned from the device
4) AVAILABLE
  Plug-In has received the latest measurement from the device and it is available for applications
5) MEASUREMENT_ERROR
  Measurement has somehow failed, e.g.; user has detached the device from body during measurement.
These states represent what are useful for applications.

## Slide 4: Measurement state transition
5) MEASUREMENT_ERROR
1) DISCONNECTED
2) READY
3) MEASURING
4) AVAILABLE
Normal flow
Error flow
Continuous measurement
(additional)

## Slide 5: Policy for one-shot response
Targeted use case of one-shot API is to:
  Retrieve the latest result of one measurement that has been completed
Definition of the policy on how a response is returned for a one-shot API request
  If the Plug-In has the latest measurement result:
    The Plug-In SHALL return the result immediately
  Otherwise:
    The Plug-In SHALL return an error with an error code that represents the measurement state of the Plug-In

## Slide 6: Types of measurement
Single measurement
  This type of devices return one result per measurement
  e.g., Thermometer, Wight Scale/Body Composition Analyzer, Glucometer, Blood Pressure
Continuous measurement
  This type of devices return one result at a time periodically for a long time
  e.g., Pulse Oximeter, Heart Rate Monitor
Issue:
  How should Plug-In respond to the app for an one-shot request?

## Slide 7: Example of single measurement (Normal case)
Measurement
Result
  e.g., Thermometer, Wight Scale/Body Composition Analyzer, Glucometer, Blood Pressure
One-shot Request
Plug-In SHALL return an error with error code 1
One-shot Request
Plug-In SHALL return an error with error code 2
One-shot Request
Plug-In SHALL return an error with error code 3
One-shot Request
Plug-In SHALL return the latest result with error code 0
State of Plug-In
1) DISCONNECTED
2) READY
3) MEASURING
4) AVAILABLE
Application

## Slide 8: Example of continuous measurement (Normal case)
4) AVAILABLE
4) AVAILABLE
4) AVAILABLE
  e.g., Pulse Oximeter, Heart Rate Monitor
One-shot Request
Plug-In SHALL return the latest result with error code 0
One-shot Request
Plug-In SHALL return the latest result with error code 0
State of Plug-In
Measurement
Result
Measurement
Result
One-shot Request
Plug-In SHALL return an error with error code 1, 2, or 3
One-shot Request
Plug-In SHALL return the latest result with error code 0
Measurement
Result
1) DISCONNECTED
2) READY
3) MEASURING
Application

## Slide 9: Definition of error code
The value of "result" is able to be used as an error code
  The "result" has already been specified in the GotAPI 1.1 spec and each PCH TS
The definition of error code will be added in each PCH TS
Name | Type | Definition of value | Mandatory/ Optional
result | Int | If success, the value is 0, otherwise an integer greater than 0, which indicates an error code. 0 means that the Plug-In is in the state of "available". If the Plug-In is not able to report any measurement result and is to response an error to the application, the Plug-In SHALL set a value representing the state of the Plug-In as follows: 1: the Plug-In is in DISCONNECTED 2: the Plug-In is in READY 3: the Plug-In is in MEASURING. 4: the Plug-In is in MEASUREMENT_ERROR. | Mandatory
* The definition colored in red will be newly added.

## Slide 10: Policy for asynchronous messages
Plug-In SHALL send a measurement result to the application upon Plug-In obtaining a new measurement result from the device
If Plug-In does not obtain the measurement result, Plug-In SHALL send an error with an error code that represents the measurement state of the Plug-In
This policy will be reflected to all the PCH TSes